DEVICE UPDATE

11/11/2024

To:	Mercy and St. Luke’s Cedar Rapids, Affiliate Sites Radiology Department Managers and MRI Technologists

From:	Hannah Monahan MD, MRMD

Date:	11/11/2024

Update:

The Abbott FreeStyle Libre 2, FreeStyle Libre 2 Plus and FreeStyle Libre 3 Plus continuous glucose monitoring (CGM) systems have received FDA clearance for wear during magnetic resonance imaging if specific imaging conditions are met.

Exclusions: This update does not apply to FreeStyle Libre Pro or FreeStyle Libre 14-day continuous glucose monitoring systems.

MRI conditions (additional information attached):

- 1.5T scanner: Up to 1 hour of continuous scanning without a cooling period, no body part restriction

- 3.0T scanner:

  - Between sternum and pelvis: Up to 12 minutes of scanning with a 2-minute cooling period between scans.

    - Musculoskeletal protocols

      - Arthrograms: hip and shoulder

      - Upper extremity: elbow, wrist, hand

      - Sacroiliac joints

      - Sternoclavicular joints

    - Abdominal protocols

      - Perianal fistula

      - Female pelvis

      - Anorectal cancer

      - Prostate

  - Anatomy outside of #2a: Up to 1 hour of continuous scanning without a cooling period.

Given the current demand for MR imaging, time constraints and required cooling period, the CGMs specified above MUST BE REMOVED prior to any 3.0T exams performed between the sternum and pelvis.

It is important to note that device readings may be inaccurate during active MRI scanning and up to 1 hour following MRI exposure. Patients should be made aware of the potential for false glucose readings prior to and after their MRI examination.

Specific exam recommendations per RCI:

- Musculoskeletal

  - Shoulder, hip

    - Arthrogram: Remove CGM, scan at 3.0T

    - Non-arthrogram: Do not remove CGM, scan at 1.5T

  - Elbow, wrist, hand

    - Positioned prone with arm above head: Do not remove CGM, scan at 3.0T

    - Positioned supine with arm by side: Can scan at 1.5T if the patient does not want to remove CGM. Scan at 3.0T if removed

  - Knee, ankle, foot: Do not remove CGM, scan at 1.5T or 3.0T per protocol

  - Sacroiliac joints: Can scan at 1.5T if the patient does not want to remove CGM. Scan at 3.0T if removed

  - Sternoclavicular joints: Do not remove CGM, scan at 1.5T

- Neuroradiology

  - Spine: Do not remove CGM, scan at 1.5T

    - If a full spine cannot be completed in less than an hour the exam must be performed on separate dates or the CGM must be removed.

  - Brain: Do not remove CGM, scan at 1.5T or 3.0T per protocol

- Abdomen

  - Liver, pancreas, MRCP, adrenal, kidney: Do not remove CGM, scan at 1.5T

  - Enterography, urogram: Do not remove CGM, scan at 1.5T

  - Perianal fistula:  Can scan at 1.5T if the patient does not want to remove CGM. Scan at 3.0T if removed

  - Female pelvis: Can scan at 1.5T if the patient does not want to remove CGM. Scan at 3.0T if removed

  - Anorectal cancer, prostate: Remove CGM, scan at 3.0T

Sources:

FDA 510(k); accessed June 2024 from https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm?ID=K233537

Matievich, William, Namvar Kiaie, and Timothy C. Dunn. "Safety and Functional Integrity of CGM Sensors When Used During Radiologic Procedures Under High Exposure Conditions." Journal of Diabetes Science and Technology (2022). https://doi.org/10.1177/19322968221106206